## Slide 1: MUX Circuit for Electrical Impedance Tomography (EIT) System
Weekly report – W1

## Slide 2: System overview
Fig 1. Block diagram of Electrical impedance tomography (EIT) system (5)
Notes: Mainly consists of 3 main parts: Acquisition system, phantom, host system
Acquisition system includes an MCU, an impedance measurement module and multiplexer circuit.

## Slide 3: Drive pattern: 4-wire measurement using AD5941

## Slide 4: Multiplexer (MUX) circuit
Controlled by Arduino Mega (1)
74HC4067 – 8 channel
Controlled by Raspberry Pi (2)
ADG506AKNZ – 16 channel
1ms/data -> 4s/measurement
Controlled by Arduino Mega (4)
ADG506AKRZ – 16 channel
Switching time: <1us
-> Require 4 MUX as analog switches
Notes: 100us

## Slide 5: 4 Current injection methods - Neighboring
Red: current electrodes
Blue: voltage electrodes
208 voltage measurement
104 independent data
the differential potentials are measured across the different electrodes, excluding the current electrodes

## Slide 6: 4 Current injection methods - Opposite
Red: current electrodes
Blue: voltage electrodes
208 voltage measurement
104 independent data
Differential potentials are measured on the voltage electrodes with respect to the electrode (called as the voltage reference electrode) adjacent to the current-injecting electrode.
current distribution is more uniform and, therefore, has a good sensitiviy

## Slide 7: 4 Current injection methods - Cross
Red: current electrodes
Blue: voltage electrodes
182 voltage measurement
104 independent data
Does not have as good a sensitivity in the periphery as does the neighboring method, but has better sensitivity over the entire domain under test

## Slide 8: 4 Current injection methods - Adaptive
Red: current electrodes
Blue: voltage electrodes
102 independent data
the boundary potentials are measured with respect to a single grounded electrode

## Slide 9: Future planning
Research more about the speed for data acquisition -> to select MUX and MCU to control the MUX

## Slide 10: Reference
https://www.sciencedirect.com/science/article/pii/S2468067222000190
(PDF) Design of low-cost and high-speed portable two-dimensional electrical impedance tomography (EIT) (researchgate.net)
Studying the resistivity imaging of chicken tissue phantoms with different current patterns in Electrical Impedance Tomography (EIT) – ScienceDirect
An open-source and easily replicable hardware for Electrical Impedance Tomography – ScienceDirect
Deep learning scheme PSPNet for electrical impedance tomography (researchgate.net)

## Slide 11: Microcontroller - Requirement
Have enough GPIO pins to control 4 muxes
  Each mux requires 4 control pins (INH, A, B, C) -> Require at least 16 pins + 2 pins for I2C communication
Have I2C communication
Low-cost
Notes: Toc do i2c